# TECHNICAL DATA SHEET

Product Name: Personal Intelligence Measurement

Patent Pending: Systems and Methods for Psychological Monitoring (US 19/227,299)

## Product Overview

The Real-Time Embedded Emotional Dimension Tracker is a lightweight, MCU-optimized implementation of a patented method for inferring emotional state from PPG/IBI data. It outputs three validated dimensions — Valence, Motivation, and Activation — in real time, enabling wearable, clinical, and HMI devices to integrate advanced affective sensing without cloud or phone processing. These dimensions allow for high level discrete emotion detection (e.g., fear, anger, sadness, joy, contentment) with training data only optional, mood detection, and personality measurement. Potential applications of these real time (both immediately measured and high temporal accuracy) measures are abundant across domains (e.g., fitness, sleep tracking, wellbeing, mental health, medical, workplace).

### Key Benefits:

- Real-time operation on wearable-class microcontrollers

- Validated against physiological and behavioral gold standards

- Compact C implementation using single-precision floats

- Flexible input: raw PPG or pre-computed IBI and optionally PVA

- Minimal code and RAM footprint

## Core Capabilities

| Feature | Specification / Description |
| --- | --- |
| Measured Dimensions | Valence (PNS), Motivation (α-SNS), Activation (β-SNS) |
| Input Signals | Raw PPG ≥ 64 Hz or IBI |
| Output Rate | Per beat or 1–5 Hz |
| Validation | Correlated with EAR, EMG, and joystick measures |

## Hardware Requirements

| Parameter | Minimal Requirements |
| --- | --- |
| Processor | Minimum: ARM Cortex-M4 or M33 @ ≥64 MHz with FPU Recommended: ARM Cortex-M4/M7/M33 @ 100–150 MHz with FPU |
| RAM | Minimum: 64 KB RAM Recommended: 128 KB+ |
| Flash | Minimum: 128 KB Recommended: 256 KB+ |
| Optional Peripherals | I²C/SPI PPG sensor, BLE module |

### Constrained / Lite Options

The full Emotion Engine runs both IBI and PVA feature chains with independent EKFs, but there are multiple ways to scale down requirements depending on the hardware budget:

- Drop PVA Features – Disable PPG amplitude features and their EKFs. Keeps only IBI-derived HRV (time + frequency) for Valence, Motivation, Activation.

- Reduce Raw PPG Sampling Rate – Beat detection normally expects high sampling (e.g., 2 kHz), but can tolerate lower rates (250–500 Hz) with minimal accuracy loss, cutting both CPU and memory load.

- Skip PPG Wave Normalization – Removes a preprocessing stage; reduces buffer usage and multiplications. Recommended for constrained devices where amplitude fidelity is less critical.

- On-Board Beat Detection – If the device already provides RR/IBI timestamps, the library can bypass raw PPG altogether and use the device’s detections. This does not lower requirements, but it can lean on existing hardware pipelines.

With all Lite optimizations applied (IBI-only + lower sampling + no normalization), the system can run in ~32 KB RAM and <120 KB flash, fitting even on smart rings and low-power Cortex-M0+/M3 devices. Accuracy increases incrementally as features are re-enabled on more capable cores (M4F/M33F+).

## I/O Requirements

- Input: Photoplethysmography (PPG) data stream

  - Sample rate: 64–2000 Hz

  - Resolution: ≥12-bit ADC or equivalent digital output

  - Interface: I²C (≥400 kHz) or SPI (≥1 MHz)

- Optional Input: Beat-to-beat interval data (if using IBI-only mode)

- Output: Floating-point values of Valence, Motivation, Activation (last beat updated per beat)

  - Latency: Typically <50 ms after beat detection (device dependent)

## Performance Metrics

Latency & CPU Load (Estimated from PC)

| Mode | Avg Latency (µs) | p95 Latency (µs) | CPU % Avg | CPU % p95 |
| --- | --- | --- | --- | --- |
| Kernel (Features → Dimensions) | 37.5 | 37.5 | 0.0047% | 0.0047% |
| Full Pipeline (From PPG Wave) | 108.7 | 103.2 | 0.0135% | 0.0128% |

Interpretation: Even the full end-to-end pipeline costs ~109 µs per beat. At normal heart rates this is ~0.0135% CPU, rising only to ~0.022% CPU at 120 bpm—still negligible.

## Integration & Deployment

- Interfaces: C API with defined input/output structures

- Dependencies: None (self-contained implementation)

- Licensing: Evaluation license; production license (per-unit or royalty)

- Integration Support: Sample code, golden (validation) test sets

## Validation Evidence

Results from FDA functional correlation with circular permutation tests with validation measures.

### Full Model:

| Dimension | Correlation (r) | Functional r | p | Validator |
| --- | --- | --- | --- | --- |
| Valence | 0.21 | 0.20 | 0.018 | EMG |
| Motivation | 0.22 | 0.12 | 0.058 | Joystick |
| Activation | 0.26 | 0.14 | 0.008 | EAR |

### Test Conditions:

A human-subjects validation study was conducted at Virginia Tech (IRB-approved) to evaluate the patented algorithm’s ability to predict three continuous emotional dimensions — Valence, Motivation, and Activation — from photoplethysmography (PPG) at every heartbeat. The dimensions were validated by EMG (zygomaticus minus corrugator) for valence, a joystick for motivation, and eye aspect ratio (EAR) for activation.

Graphs and plots available upon request.

## Contact Information

Contact Name: Shane Ryan McClafferty

Email: srmcclafferty@gmail.com

Demo Availability: Software trials available through C library files (.a and .h files).

Customization: Adaptation for other sensors or applications on request. Potentially requiring research or data collection. Open to agreements for further research.